## Slide 1: Harmonization of Friction Measuring Devices Using Robust Regression Methods
Samer Katicha
09/09/2013

## Slide 2: Outline
What is harmonization of devices?
  Measurement conversion
  Conversion accuracy
The best case scenario
  Perfectly identical devices
Failure of Linear regression
  Ideal (almost utopian) conditions
  Playing a trick on you
  Biased and inconsistent estimate
Notes: 1 slide only

## Slide 3: Outline
Measurement error models
  Orthogonal regression
  The general model
Practical results
  Repeated measurements from the same device (locked wheel tester)
  Simulation of two devices
  Measurements from two different devices (locked wheel testers)
Notes: 1 slide only

## Slide 4: Harmonization
ASTM
  “the adjustment of the outputs of different devices used for the measurement of a specific phenomenon so that all devices report the same value”
Does not consider agreement
  Confidence between adjusted outputs
  Limits of agreement (LOA)
Notes: 1 slide only

## Slide 5: Perfectly Identical Devices
This is the best we can hope for
It means:
  If we have two devices and obtain two sets of measurements (runs) from device 1, this is the same as:
  Obtaining two sets of measurements from device 2, or
  Obtaining two set of measurements each from one of the devices
Notes: 1 slide only

## Slide 6: Failure of Linear Regression
Notes: 1 slide only

## Slide 7: Ideal Conditions
Source of errors restricted to independent identically distributed normal error
The measurements
  yi = measurement (with error)
  fi = true value
  i = error which is N(0,2)
Notes: 1 slide only

## Slide 8: Playing a Trick on You
You own two locked wheel skid testers
You ask the operator to test the same pavement sections with the two devices (1 test each)
The operator, being lazy, performs the two tests with device 1
Testing conditions were ideal: nothing affected results other than random error
To compare the devices you do linear regression and find…
Notes: 1 slide only

## Slide 9: Biased Estimate
Notes: 1 slide only

## Slide 10: Inconsistent Estimate
Notes: 1 slide only

## Slide 11: Reasons
Ordinary regression assumes:
  Only one of the variables contains error
  x variable is measured “without” error
When both variables contain error:
  Ordinary regression (least squares) is Biased
  Bias depends on amount of error compared to range of measurement
  More error more bias
Notes: 1 slide only

## Slide 12: Ordinary Regression Summary
Biased estimate
Inconsistent estimate
This occurs in the most simple case:
  Error normally distributed and no other factors affecting measurement
  Real measurements are more complicated
Things can get worse:
  Speed correction models not perfect: introduce MORE error
  Texture models introduce MORE error
Notes: 1 slide only

## Slide 13: Measurement Error Model
Notes: 1 slide only

## Slide 14: Measurement Error Model
Take into account errors in both variables
Simplest case: orthogonal regression
Gives an “unbiased” estimate
Notes: 1 slide only

## Slide 15: Orthogonal Regression
Notes: 1 slide only

## Slide 16: Orthogonal Regression
Notes: 1 slide only

## Slide 17: Postulating the Model
The best case only considers measurement error
Real measurements:
  Measurement error (repeatability)
  Model/device error:
    device estimates the “true” friction
    Independent of measurement error
    Experimental evidence of device error: between devices error
Notes: 1 slide only

## Slide 18: Postulating the Model
    Y = device estimate (with no repeatability error)
    X = “true” friction which is not known
    ay and by = model parameters
    Y = model error
     = device measurement
    rY = repeatability error
Notes: 1 slide only

## Slide 19: Postulating the Model
    Y = device estimate (with no repeatability error)
    X = “true” friction which is not known
    ay and by = model parameters
    Y = model error
     = device measurement
    rY = repeatability error
Notes: 1 slide only

## Slide 20: Harmonization
Notes: 1 slide only

## Slide 21: Harmonization
     = repeatability error standard deviation
     = model error standard deviation
    Solution is found iteratively
Notes: 1 slide only

## Slide 22: Practical Results
Notes: 1 slide only

## Slide 23: Measurements from Same Device
Notes: 1 slide only

## Slide 24: Model Parameters
True parameters:
  Slope = 1
  Intercept = 0
Ordinary regression
  Slope = 0.9262
  Intercept = 3.3002
Measurement error model:
  Slope = 1.0207
  Intercept = -2.1425
Notes: 1 slide only

## Slide 25: Dangers of Speed and Texture Correction
Notes: 1 slide only

## Slide 26: Model Parameters
True parameters:
  Slope = 1
  Intercept = 0
Ordinary regression
  Slope = 0.7905 (previous 0.9262)
  Intercept = 10.1880 (previous 3.3002)
Measurement error model:
  Slope = 1.0448 (previous 1.0207)
  Intercept = -4.5637 (previous -2.1425)
Notes: 1 slide only

## Slide 27: Different Devices (Simulation)
Notes: 1 slide only

## Slide 28: Different Devices (Simulation)
Notes: 1 slide only

## Slide 29: Different Devices (Simulation)
Notes: 1 slide only

## Slide 30: Different Devices (Simulation)
True parameters:
Ordinary regression
Measurement error model:
Notes: 1 slide only

## Slide 31: Measurements from Different Devices (Locked Wheel)
Notes: 1 slide only

## Slide 32: Model Parameters
Averaged:
  Ordinary regression:
    Slope = 0.8927; Intercept = 10.3757
  Measurement error model:
    Slope = 0.9341;Intercept = 8.3281
Not Averaged:
  Ordinary regression:
    Slope = 0.8590 (bias = 0.8927-0.8590 = 0.0337)
  Measurement error model:
    Slope = 0.9346 (bias = 0.9341-0.9346 = -0.0005)
Notes: 1 slide only

## Slide 33: Limits of Agreement (LOA)
Relationship between devices:
  Convert between measurements
  Does NOT give agreement
After conversion:
  Limits of Agreement (LOA)
  General: Bland and Altman (1983), Altman and Bland (1986)
  Friction devices: De Leon et al. (2012)
Notes: 1 slide only

## Slide 34: Conclusions
Ordinary regression:
  Biased estimate of the relationship
  Inconsistent relationship based on choice of dependent and independent variable
Measurement Error Model:
  Unbiased estimate of the relationship
  Relationship independent of choice of dependent and independent variables
Notes: 1 slide only

## Slide 35: Conclusion
ASTM
  “the adjustment of the outputs of different devices used for the measurement of a specific phenomenon so that all devices report the same value”
Ordinary regression
  The same device WILL NOT report the same value
Notes: 1 slide only

## Slide 36: Reproducible Research
Paper has been submitted to TRB
Paper written under the reproducible research paradigm:
  All data used in the paper is made available for anybody to use
  Matlab files used are made available:
    Calculations organized in functions
    GUI that can be used to reproduce all Figures showing the results
Notes: 1 slide only

## Slide 37
Blacksburg, VA